	





TECHNICAL REPORT

19902 Altaire ll Anti-decubitus mattress system
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1. PRODUCT DESCRIPTION

1.1 Intended use
19902 ALTAIRE II is a dynamic anti-decubitus mattress system using an air mattress made from nylon & polyeurathane and powered by an air pump to redistribute pressure across the surface of the mattress to avoid vascular compression particularly on areas of bony protrusion such as heels or the sacrum . It is a dynamic surface indicated for the prevention of decubitus ulcers up to stage III and it is available for patients up to 180 Kg. The particular structure of the mattress, its modularity and flexibility make it adaptable to all kind of profiling beds in a domestic or hospital environment.
To use properly the mattress, we suggest to call on competent or paramedical/medical staff. According to D. Lgs. 81/2008, the surface can be managed in complete safety.

1.2 System description
19902 ALTAIRE II mattress system (Fig. 1) is composed by a dynamic anti-decubitus mattress with its cover and a pump that powers the mattress. The mattress can be positioned directly on the bed structure to replace the traditional mattress and it can also fit electrical adjustable bed movements. Thanks to the serigraphy printed on the cover, it is very easy to recognize the upper side of the mattress 19902 ALTAIRE II and unlike traditional mattresses, 19902 ALTAIRE II must not be periodically flipped. The Air pump ALTAIRE II has an internal air compressor to deliver 8 litres of an air per minute through the two air ports on the side of the pump to inflate the mattress.


Figura 1. 19902 ALTAIRE II



1.2.1 Antidecubitus mattress
The 19902 ALTAIRE II anti-decubitus mattress is made of three different layers that perfectly adapt to the body shapes and guarantee lower values of contact pressures while providing a high comfort level.
The first bottom layer is the mattress base made from Nylon and holds within its walls and base the second and dynmaic part of the mattress.The second part is a series of up to 17 cells made from polyeurythene that alternately fill with air to provide support but then after 5 minutes empty out to give pressure relief whilst the cell next to it fills with air to give support. The first cell from the foot end of the mattress and each alternate cell are known as A cells and the cells inbetween are known as B cells with A & B cells alternately giving support or relief. The modular structure of the cells mean that the mattress profiles to most hospital & nursing beds. The third and last layer is the cover which has high stretch properties to help it conform to patient movements
The bearing structure is made of PU. All Polyurethane material used are Latex-free, Phthalate-free and fire retardant.
The system is totally radiotransparent.


1.2.2 Cover
The upper cover is made of Polyester-Polyurethane .This particular material used is bi-elastic (with a low friction coefficient), waterproof and resistant to body fluids, vapour perspirant, fire resistant, smell and spot repellent, anti-bacterial, anti-fungal, antistatic, Latex-free and Phthalate-free.
The entire cover is easy to be removed from the mattress and it is interchangeable through a zip, which is properly covered by a flap; The cover is easily cleanable .Clinical tests on the upper cover confirm that the material doesn’t produce any kind of allergic reaction. Further details are reported in the data sheet section. Tests also show that the cover is completely non-toxic for patient’s skin and oral tract.

1.2.3 Pump

The ALTAIRE II pump is a Medical Device Class 1 when used in conjuction with mattress 19902 ALTAIRE II . It is a 220V-240V 50HZ mains powered pump using a quiet internal air compressor to alternatively deliver air to the mattress 19902 ALTAIRE II A cells and then by use of the internal rotor valve it will then 5 minutes later deliver air to mattress 19902 ALTAIRE II B cells using the two A&B air ports on the side of the pump. Use of intelligent pressure sensing technology means that the pump switches off when the optimum internal pressure is reached. Weight of the patient is entered using the dial to correctly reach the best internal pressure.

	Power consumption
	7W (Altaire ll 7W)

	Fuse rating
	F2A 250VAC

	Electrical isolation
	Class I 2a

	Ingress protection
	IP21

	IEC conformity
	60601-1, 60601-1-2, 60601-11

	Warranty
	2 years against all manufacturing faults




2. DATASHEET



	Manufacturer
	Sequoia Healthcare

	Production date
	2021

	Device
	Anti-decubitus dynamic mattress system

	Mattress Type
	Dymanic mattress system (powered)

	Classification, according to
93/42/CEE
	Medical device Class I

	Intended use
	Prevention of decubitus ulcers up to stage III

	Conformity
	93/42/CEE and subsequent modifications and additions UNI EN ISO 13485 and UNI EN ISO 9001
Medical Devices Directive 2007/47/EC

	Warranty
	2 years (for both the mattress, pump )

	
	




2.1. Mattress specifications



	System description
	Anti-decubitus dynamic mattress system (powered by its supporting air pump) mattress made from PU cells to support the patient and alternatively give support and pressure relief and encased in Nylon base & PU zipped cover
Latex and Phthalate free. Completely radiotransparent device.

	Mattress standard size
	198 x 88 x 12cm (non-standard size on request)

	Device’s weight
	6 Kg

	
Maximum supported weight
	
150 Kg

	Fire resistance
	Component materials (Polyurethane and the entire cover) are certified by the manufacturer as fire resistant







2.2.1 Cover specifications

	
Composition
	Upper cover: Polyester-Polyurethane; it is antifungal, antibacterial, breathable to water steam, waterproof and resistant to liquids, smell and
spot repellent, fire resistant and machine washable. Latex and Phthalate free.

	MECHANICAL AND PHYSICAL PROPERTIES OF UPPER COVER
	STANDARD

	Waterproofness
	> 3000 mm
	EN ISO 811

	Breathability
	>= 600 g/m2 x 24 h at 38°
	ASTM E96


	Fire standard
	CRIB 5
	BS 6807- BS 7175

	
Notes
	· Permanent antimycotic and antibacterial barrier; moisture resistant
· Resistant to all organic fluids
· Washing prescriptions are printed on the cover













1.. 2.2 Pump specifications



	Pump description
	Anti-decubitus dynamic pump. Connects to the mattress via quick release connectors . Two air ports alternatively deliver air to the A cells and then after 5 minutes the B cells through the two hoses at the foot end of the mattress
.

	Pump standard size
	24cm wide x 11.5cm high x 9cm deep. 2.2kg

	Device’s weight
	2.2KG




2.1.1 Cover specifications

	
Composition
	Upper cover: Polyester-Polyurethane; it is antifungal, antibacterial, breathable to water steam, waterproof and resistant to liquids, smell and
spot repellent, fire resistant and machine washable. Latex and Phthalate free.

	MECHANICAL AND PHYSICAL PROPERTIES OF UPPER COVER
	STANDARD

	Waterproofness
	> 3000 mm
	EN ISO 811

	Breathability
	>= 600 g/m2 x 24 h at 38°
	ASTM E96

	Fire standard
	CRIB 5
	BS 6807- BS 7175

	
Notes
	· Permanent antimycotic and antibacterial barrier; moisture resistant
· Resistant to all organic fluids
· Washing prescriptions are printed on the cover



2.1. Washing prescriptions and system sanification (figure 1)


	







Cover washing prescriptions
	
· Machine washable up to 70°C(*) with a proper non aggressive cleaning agent

· The normal cleaning instructions are to wipe the surface clean with a soft sponge or soft cloth with warm soapy water



The upper cover can tolerate Chlorine releasing agents at maximum 1,000 Parts Per Million (0.1% ppm); whenever Chlorine solution is used, the cover must be washed afterwards with water so the chlorine residue does not stay on it. However, please be aware that the stronger is Chlorine solution used (max 0.1% ppm) the higher rate of upper cover replacement will be required in time.

	Disinfection
	The cover can be sanitized in autoclave at a temperature of 134°C for 3 minutes; please first clean the surface with pure water



(*) A temperature of 60°C is suggested to perform an effective disinfection while better preserving the serigraphy and the material over time.



2.3 Specifications

	Technical Specifications
	
	Atmospheric Specifications

	Therapy modes
	Dynamic & Static
	Storage temperature
	-25 - 70°C

	Compressor air flow
	Approx. 8lpm
	Transport temperature
	-25 - 70°C

	Cycle time
	10 minutes
	Operating temperature
	5 - 40°C

	Anti-particle filter
	Yes
	Humidity
	10 - 90%

	Visual and audible
alarms
	Yes
	Atmospheric pressure
	700 - 1060hPa

	Electrical power supply
	
	220 – 240V / 50Hz
	

	Power consumption
	
	7W (Airwave 1010: 8W)
	

	Fuse rating
	
	F2A 250VAC
	

	Electrical isolation
	
	Class I 2a
	

	Ingress protection
	
	IP21
	

	IEC conformity
	
	60601-1, 60601-1-2, 60601-11

	Warranty
	
	2 years against all manufacturing faults



	CE marking

	CE Marking indicating conformance to EC Directive No. 2007/47/EC concerning medical devices

	Type BF Applied Part (patient isolation from electrical shock)

	Class II Product
Indicates separate collection for electrical and electronic equipment (WEEE) IP Protection rating: IP21 – protection against finger and dripping water
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